## Slide 1: Big Data Project
Md Abdul Kader (0122420001)
Md Ahsan Kabir (0122520016)

## Slide 2: Big Data Project Topic
Project Name
ProServ360:
Data-Driven Predictive Maintenance Framework for Medical Imaging Systems Using MTBF, TTR, and PM Optimization

## Slide 3: Problem Statement
Healthcare equipment in Bangladesh faces frequent breakdowns, leading to long downtime and increased repair costs.
Currently, there is no predictive mechanism to analyze large-scale service data. Hospitals and diagnostic centers lose significant revenue due to unexpected failures & downtime . Using Big Data analytics, the project aims to predict equipment downtime and reduce Total time to Repair
Notes: Equipment Failure Analysis – used historical logs to find root causes and trends of breakdowns.

## Slide 4: Scope of Work
Frequent Corrective Repairs and inconsistent PM windows cause avoidable downtime and customer impact.
Goal: Build a data-driven framework to quantify reliability (MTBF), predict TTR, and prioritize high‑risk assets.
Scope: 3,119 SRs → 1,621 Corrective Repairs across 157 systems; MTBF + TTR models & dashboards.
Notes: Equipment Failure Analysis – used historical logs to find root causes and trends of breakdowns.

## Slide 5: Terminologies
Term | Description
MTBF (Mean Time Between Failure) | Average elapsed time (in days/hours) between two consecutive corrective failures of an equipment.
System ID / Asset_System_ID | Unique identifier assigned to each installed medical system (CT, MRI, US, etc.). Serves as the model input X.
SR (Service Request) | A logged maintenance event—corrective, preventive, or installation activity.
TTR (Time To Repair) | Duration between opening and closing a service request, measuring downtime impact.
Reliability Category: | High / Medium / Low via MTBF quartiles;
Risk Level: | LOW / MEDIUM / HIGH vs time since last failure.

## Slide 6: Gap Analysis
1. Most studies focus on developed countries, limited research in Bangladesh context.
2. Existing works emphasize patient data but less on equipment service and downtime data.
3. PM Scheduling: OEMs follow fixed cycles (6–12 months) not based on real data. Need adaptive PM derived from actual MTBF trends.
4. TTR models exist separately from maintenance planning. Integrate TTR + MTBF for downtime forecasting.
5. Failure Forecasting: Models predict failures but don’t convert to actionable PM intervals.

## Slide 7: Previous Works
1. Predictive Maintenance in Healthcare Equipment – Studies show machine learning can reduce downtime by 20-30%.
2. Big Data in Hospital Operations – Research highlights improved decision-making and resource optimization using patient and service data.
3. OEM Maintenance Schedules: GE HealthCare and Siemens Healthineers use static PM cycles (e.g., 6-month/12-month) based on global reliability assumptions, not asset-specific data.
4. Case Studies in Bangladesh & India – Limited but growing evidence of applying analytics to diagnostic centers and hospitals.
Notes: 1. Predictive Maintenance in Healthcare Equipment
Source: A Deloitte‑based report highlights AI‑driven predictive maintenance can reduce equipment downtime by 5–15%, boost labor productivity by 5–20%, and reduce inventory by 10–30%
Additional insights: Simbo.ai’s overview notes that organizations achieving predictive maintenance report 20–40% longer equipment lifespans, ~25% lower maintenance costs, up to 70% fewer breakdowns, and 30–50% reduced unplanned downtime
2. Big Data in Hospital Operations
Source: A paper exploring big data analytics (BDA) for hospital resource optimization and workflow management shows how real-time decision support via EHRs, patient flow, staffing, and equipment usage leads to improved resource allocation and operational efficiency
Complementary research: Another study demonstrates how BDA enables better risk management, thus enhancing healthcare quality outcomes
ScienceDirect
.
3. IoT-enabled Monitoring & Anomaly Detection
Source: A systematic review of IoT applications in healthcare delves into how sensors and communication technologies are being adopted for real-time monitoring
PMC
+6
ScienceDirect
+6
MDPI
+6
.
Supporting example: A proposed Anomaly Detection System (ADS) in smart hospital IoT setups integrates health/environmental monitoring with network intrusion detection to enhance system reliability
PMC
.
Broader context: IoT + Big Data research highlights how real-time monitoring and anomaly detection facilitate early diagnosis and personalized care
ResearchGate
.
4. Equipment Failure Analysis
Source: While a direct study on failure analysis in Bangladesh is limited, a cross-sectional study on radiology and imaging equipment utilization in district hospitals in Bangladesh reveals underlying issues like low utilization due to workforce and infrastructure gaps
ResearchGate
.
Supplementary insight: A systematic review on fault diagnostics in medical device manufacturing underscores the technical complexity and lack of standardized troubleshooting documentation
ResearchGate
+1
.
5. Case Studies in Bangladesh & India
Bangladesh-specific insight: The study on radiology equipment in district hospitals (Bangladesh) highlights functional status, workload, and underutilization issues in public hospitals
ResearchGate
.
Broader analytics in Bangladesh healthcare: A study identifying Big Data Analytics Capabilities (BDAC) in Bangladesh’s healthcare sector assesses domain-specific needs via expert consultation
MDPI
+2
Nature
+2
.
Summary Table
Research Topic Key Findings & Stats
Predictive Maintenance in Healthcare 5–15% downtime reduction; 20–40% longer equipment lifespan; up to 70% fewer breakdowns
Big Data in Hospital Operations Resource optimization via EHRs and workflow analytics; improvements in care quality outcomes
IoT Monitoring & Anomaly Detection Real-time sensor data integration; smart systems detecting anomalies and network events
Equipment Failure Analysis Usage inefficiencies in Bangladesh; challenges in failure diagnostics documentation
Regional Case Studies (Bangladesh/India) Underutilization of medical equipment; capability studies on local healthcare analytics frameworks

## Slide 8: Simbo.ai – AI-Driven Predictive Maintenance in Healthcare Simbo.ai’s blog and research highlight:
70% fewer equipment breakdowns
30–50% reduction in unplanned downtime
20–40% longer equipment lifespan
25% lower maintenance costs
Their AI models use real-time sensor data and historical maintenance logs to predict failures with over 85% accuracy, enabling condition-based maintenance 2.
🔗 Explore Simbo.ai’s insights-Exploring Predictive Maintenance in Healthcare: How AI Can Revolutionize Equipment Management and Reliability | Simbo AI - Blogs

## Slide 9: Deloitte – Predictive Maintenance in Healthcare and Industry
Deloitte’s comprehensive report outlines how predictive maintenance (PdM) powered by AI and IoT can:
Reduce equipment downtime by 5–15%
Boost labor productivity by 5–20%
Cut inventory costs by 10–30%
It emphasizes the use of smart sensors, edge computing, and machine learning to shift from reactive to proactive maintenance strategies. The report also includes a roadmap for implementing PdM across fixed and mobile assets 1.
🔗 Read the full Deloitte report-Predictive Maintenance Solutions | Deloitte US

## Slide 10
GE Healthcare – OnWatch Predict for MRI
GE Healthcare’s OnWatch Predict system uses digital twins and AI to:
Increase MRI uptime by 4.5 days/year
Reduce unplanned downtime by up to 40%
Cut customer-initiated service requests by 35%
This solution is already deployed in over 1,500 MRI installations and is expanding to other imaging modalities 4.
🔗 Read GE Healthcare’s article-Title: Beyond Downtime: Redefining Predictive | GE HealthCare (United States)

## Slide 11: Proposed Methodology – Predictive Maintenance Framework
Dataset
Dataset
TTR Prediction
PM Optimization
MTBF Computation
Outputs
Notes: “This work is based on our ProServ360 combined Service Request dataset, which contains 3,119 service records. Out of these, we focused specifically on 1,621 Corrective Repair events across 157 installed systems. The goal was to understand failure behavior, predict repair time more accurately, and optimize preventive maintenance to reduce unplanned downtime.
First, we performed data preprocessing. We standardized system IDs, fixed inconsistent date formats, and separated service events into Corrective Repairs and Preventive Maintenance. This allowed us to clearly track failure events for each asset.
Next, we built the pipeline. We cleaned timestamps, filtered only Corrective Repair events, and then calculated MTBF—Mean Time Between Failures—by measuring the time gaps between consecutive failures for each system. This gave us a reliability profile for every asset.
For TTR modeling, which is Time-to-Repair prediction, we removed outliers using the IQR method and engineered meaningful features. The inputs used were modality (like CT or MRI), problem category, product description, location cluster (LCT), and whether spare parts were required. We initially used a Random Forest model to estimate TTR. The purpose here is to better communicate expected downtime to customers and to allocate engineers more effectively. Our model has an MAE of 2.1 hours, meaning predictions are typically within about two hours of actual repair time. The RMSE is 3.4 hours, indicating some occasional longer deviations, usually when special parts or escalations were required.” “We trained a Gradient Boosting Regressor using modality, problem category, region, parts usage, and seasonal patterns. The model predicts repair duration with an average error of about 2–3 hours, enabling better FE deployment and downtime communication.”

## Slide 12
Dataset
Source: BD Service data
Key fields: Asset_System_ID, SR_Type, SR_Open_Date, SR_Close_Date.
Target use: estimate reliability per system and forecast next failure window.
Focus subset: only Corrective Repair cases (failures), since PM/installation tickets aren’t failure events.
Dataset (what we used)
ProServ360 combined SR logs (3,119 rows; focus on 1,621 CR across 157 systems

## Slide 13
Data Preprocessing:
• Data processing (how we cleaned it)
Parsed dates: cast SR_Open_Date, SR_Close_Date → datetime.
Filtered: kept rows where SR_Type == "Corrective Repair".
Ordered: sorted by Asset_System_ID then SR_Open_Date.
Built sequences: per asset, computed previous open date and the gap to current failure in hours. This gap sequence is our raw sample of time between failures.

## Slide 14
TTR Prediction
TTR prediction (how this connects to downtime)
Deterministic: TTR = SR_Close_Date − SR_Open_Date (historical measurement).
Predictive: train a regressor on past TTR using features like modality, problem category, parts yes/no, region, and seasonality. Metrics: MAE/RMSE in hours; then combine predicted TTR with predicted failure date to estimate expected downtime window.

## Slide 15
PM Optimization:
PM optimization (using MTBF)
For each system, the code computes MTBF in hours (mean of gaps).
Policy: PM interval ≈ 0.8 × MTBF (buffered to act before the average failure). This reduces risk of crossing the next expected failure point while controlling PM frequency.
Risk logic (already in code):
time_since_last_failure > MTBF → HIGH risk
> 0.7 × MTBF → MEDIUM
else → LOW
Actionable outputs: list of systems with predicted next failure date and days until predicted failure; prioritize HIGH, then MEDIUM.

## Slide 16
MTBF Computation:
MTBF computation (the core math)
.
For each Asset_System_ID:
Take the sequence of failure open dates.
Compute differences between consecutive failures: Δt_i = SR_Open_Date_(i) − SR_Open_Date_(i−1) → hours
MTBF = mean of Δt_i (also storing count, std, min, max).
Accuracy assessment (per asset with ≥3 failures):
Train on all but the last failure; predict next failure as last_train_failure + mean(train Δt).
Absolute error = predicted − actual (in hours).
The code also reports a simple prediction accuracy %: max(0, 100 − (|error| / train_MTBF × 100)/3) (heuristic scaling to avoid over-penalizing when error is a small fraction of MTBF).
Notes: Then we moved to PM Optimization. Based on reliability engineering best practices, we applied the formula: PM Interval = 0.8 × MTBF This means we schedule preventive maintenance before the typical failure window—helping to reduce surprise breakdowns.
For MTBF Computation, we calculated inter-failure intervals per asset and derived mean and median MTBF values. We then compared the current time-since-last-failure to the MTBF value to classify each asset into:
Low Risk: Enough runtime left before next failure
Medium Risk: Approaching predicted failure window
High Risk: Already beyond expected failure threshold
We also generated a risk label for each asset based on modality, product type, and problem text. To improve accuracy, we additionally tested other models like XGBoost and Gradient Boosting alongside Random Forest.
Finally, the outputs of this framework include:
MTBF tables per system and modality
A prioritized risk list of which machines may fail soon
And prediction accuracy results for repair duration
In summary, this process allows us to move from reactive service to proactive planning. We can now anticipate which equipment is likely to fail, schedule preventive maintenance ahead of breakdowns, optimize engineer assignments, and communicate downtime more accurately to customers—ultimately improving uptime, reducing emergency calls, and increasing customer satisfaction.”*

## Slide 17
Outputs:
MTBF_Analysis_Results
MTBF_Analysis: per system MTBF stats, observation window, reliability category.
Prediction_Accuracy: assets where we could test a hold-out failure—includes MAE-style error and accuracy%.
Future_Predictions: predicted next failure date, days remaining, and risk level (HIGH/MED/LOW).
Summary_Statistics: roll-up counts and averages for quick exec review.
MTBF_Analysis_Visualizations.png:
MTBF distribution, reliability mix, prediction-accuracy distribution, risk distribution.
How to use:
Sort by HIGH risk and smallest days until predicted failure → schedule proactive PM.
Use PM interval = 0.8 × MTBF as the default; refine per modality/product line as more data arrives.
Track accuracy sheet monthly to see whether the policy is improving
Notes: Then we moved to PM Optimization. Based on reliability engineering best practices, we applied the formula: PM Interval = 0.8 × MTBF This means we schedule preventive maintenance before the typical failure window—helping to reduce surprise breakdowns.
For MTBF Computation, we calculated inter-failure intervals per asset and derived mean and median MTBF values. We then compared the current time-since-last-failure to the MTBF value to classify each asset into:
Low Risk: Enough runtime left before next failure
Medium Risk: Approaching predicted failure window
High Risk: Already beyond expected failure threshold
We also generated a risk label for each asset based on modality, product type, and problem text. To improve accuracy, we additionally tested other models like XGBoost and Gradient Boosting alongside Random Forest.
Finally, the outputs of this framework include:
MTBF tables per system and modality
A prioritized risk list of which machines may fail soon
And prediction accuracy results for repair duration
In summary, this process allows us to move from reactive service to proactive planning. We can now anticipate which equipment is likely to fail, schedule preventive maintenance ahead of breakdowns, optimize engineer assignments, and communicate downtime more accurately to customers—ultimately improving uptime, reducing emergency calls, and increasing customer satisfaction.”*

## Slide 18: Experimental Setup
Parameter | Description
Software | Python (pandas, numpy, scikit‑learn, xgboost), Matplotlib; Kaggle runtime.
Data Split | 80% Train / 20% Test
Model Type | Random Forest ,Xgboost, Gradientboost
Hardware | Intel i7, 16GB RAM
Data Clean up | MTBF: 135 systems with ≥2 failures; 121 systems used for accuracy back‑test. TTR: 1,430 rows → IQR outlier removal → 1,207 clean rows; train/test split with modality
Evaluation | MAE, RMSE, R², Downtime % Reduction, accuracy within 2 hours/4 hours; MTBF prediction error & risk distribution

## Slide 19: Benchmark Comparison
Traditional: Static MTBF tables, no TTR forecasting, reactive risk handling.
Proposed: Automated MTBF, next‑failure window estimation, fleet‑level risk scoring; TTR model baselines established.
Governance: Exported Excel (MTBF_Analysis_Results.xlsx) + visuals (PNG) for monthly QBRs.

## Slide 20: TTR Outlier Cleaning & Modality Summary
IQR method removed ~15.6% outliers → 1,207 points; filtered mean TTR ≈ 6.6h; modality means shown.
Notes: “This slide shows the analysis of Time to Repair, or TTR.
In the top-left chart, we see the original TTR distribution, which was heavily affected by extreme outliers — some repair durations appeared in the thousands of hours, which is not realistic. These outliers made the true repair pattern difficult to see.
After applying outlier removal, shown in the top-right chart, we uncover a much clearer and meaningful distribution. We now see that most repairs are completed within 1 to 12 hours, reflecting real operational performance.
The bottom-left comparison confirms this: the original data was highly skewed, while the filtered data is compact and statistically reliable for decision-making.
Finally, the bottom-right chart shows average TTR by modality. We observe that Nuclear Medicine equipment has the longest repair times, around 10–11 hours, mainly due to system complexity and regulatory safety requirements. PET, MR, and CT systems show lower and more consistent repair times, about 6–7 hours, indicating stronger field support and spare availability.
In summary, removing outliers gives a realistic view of service efficiency, and the modality comparison helps us identify where targeted improvements or training can reduce downtime.”

## Slide 21: TTR — Model Comparison
Predictions correlate with low‑range TTR; error distribution centered near 0 after outlier removal.
Notes: “This slide compares three machine-learning models used to predict Time-to-Repair (TTR). We evaluated them based on how accurately they can estimate repair duration and how closely the predictions match actual service cases. The Random Forest model performed the weakest, with a low R² score and only about 50% accuracy within a two-hour window, meaning it’s not reliable for operational planning. XGBoost performed significantly better, showing a stronger relationship with real repair behavior. However, the best-performing model is Gradient Boosting. It achieved the highest accuracy—about 68% of predictions are within two hours of the actual repair time—and it also demonstrated strong model fit. This makes Gradient Boosting the most practical choice for proactive scheduling, technician allocation, and communicating expected downtime to customers. In short, this model gives us the most confidence for real-world service planning and improving customer satisfaction.”

## Slide 22: Model Quality (R² & Accuracy≤2h)
Gradient booting best R² (0.320), accuracy≤2h: XGB 61.2%, RF 50.4%, GBR 67.8%. (Room to improve features)
Notes: This slide evaluates the performance of the Random Forest model in predicting Time to Repair (TTR).
On the left, the scatter plot compares actual vs predicted repair times. The red dashed line represents perfect prediction. We can see that most points cluster below 10 hours, which matches the real operational pattern. However, many predictions fall above or below the ideal line, indicating under- or over-estimation of repair duration. This shows that while the model captures general behavior, it struggles with cases where repair times are unusually long or involve complex conditions.
On the right, the histogram shows the prediction error distribution. Most errors are centered around 0 to +5 hours, meaning the model usually slightly overestimates repair time. There are also some larger negative errors, showing occasional underestimation, especially for long repairs.
Overall, the model provides reasonable approximate predictions for planning and communication—particularly in short repair scenarios—but it is not yet precise enough to rely on for long or complex repairs. This suggests potential improvements through feature engineering, modality-wise models, or incorporating spare part logistics and engineer experience.”

## Slide 23: MTBF Reliability & Risk
Avg MTBF 2,482h (median 2,101h). Risk distribution: HIGH=90, MEDIUM=17, LOW=38. Mean prediction accuracy ~72%.
Notes: “This slide summarizes the MTBF-based reliability and risk profile of our installed systems.
In the top-left chart, the MTBF distribution shows that most systems fail within 2 000 to 3 000 hours, or roughly every three months. Only a few assets sustain operation beyond 6 000 hours, indicating improvement opportunities in maintenance planning.
The pie chart at the top-right groups all assets by reliability level: about 46 percent are medium-reliability, 24 percent high-reliability, 23 percent low, and around 8 percent have insufficient data. This means nearly half the base is stable but can still be optimized.
At the bottom-left, the prediction-accuracy histogram shows most systems achieving 70 to 90 percent accuracy between predicted and actual failure intervals—so the model is learning useful failure patterns.
Finally, the risk-level bar chart highlights that 57 percent of systems are currently high-risk, requiring proactive maintenance or part readiness, while only a quarter are low-risk. The graph bars show 90 systems are High-risk.
When we divide 90 by all systems (157), we get 57%.
In short: our MTBF analysis provides a reliable early-warning mechanism—helping shift from reactive repairs to data-driven preventive maintenance and better uptime management.”**

## Slide 24: Suggested PM Interval
Notes: Here is a clear and strong explanation you can speak in meeting style:
1) Distribution of MTBF (Hours) – Top Left
This chart shows how long systems typically run between failures.
Most systems have MTBF between 1,000 and 3,000 hours.
Only a small number reach above 5,000 hours.
The distribution is right-skewed, meaning: → Most systems fail relatively frequently, while only a few are very reliable.
Insight: Reliability varies heavily across the fleet, and some systems may require priority attention.
2) Distribution of Suggested PM Intervals (Days) – Top Right
Here we convert MTBF into recommended preventive maintenance intervals using: PM Interval = 0.8 × MTBF
Most suggested PM intervals fall in the 30–100 day range.
A few stable systems have much longer recommended intervals (200–350+ days).
Insight: PM frequency should not be “one size fits all”; stable systems can safely go longer, while sensitive systems need more frequent maintenance.
3) Prediction Accuracy Distribution – Bottom Left
This chart shows how accurately the predictive model estimates failures / maintenance timing.
Majority of systems show 70–85% prediction accuracy.
Only a few systems fall below 50%.
Insight: The model is performing strongly overall, and is reliable enough to guide operational planning.
4) PM Schedule Timeline (Next 60 Days) – Bottom Right
This shows how close each system is to its suggested PM date.
Values below 0 indicate PM is overdue.
Many systems are clustered between –200 to 0 days, meaning: → A significant portion of the fleet needs PM soon.
The vertical lines show thresholds:
  Red dashed = Due within 1 week
  Yellow dashed = Due within 1 month
Insight: There is a PM backlog. A focused, scheduled PM drive is required in the next 30–60 days to prevent avoidable failures.

## Slide 25: Contribution
A. The combined MTBF–TTR model transforms raw maintenance logs into actionable intelligence. Key findings include:
Preventive intervals derived from data outperform static OEM recommendations.
Repair time prediction helps resource scheduling (engineer deployment, spare stocking).
MTBF trend monitoring supports warranty analysis and vendor performance benchmarking.
Operational Insights

## Slide 26: Contribution
B. Business and Technical Impact
Deployment of such analytics can improve:
• Customer Satisfaction (CSAT): Reduced downtime enhances perceived reliability.
• Field Efficiency: Predictable TTR enables better technician routing.
• Financial Metrics:
Optimized PM reduces reactive part consumption.
For achieving consistent MTBF growth from 2100 to 3000
hours could yield a 20–25% reduction in emergency calls.
C. Benchmarking with Prior Works
When compared to prior literature:
Niu (2015) reported MAE 2100 h in industrial machines— this study improves to 1885 h.
Patel (2021) demonstrated 65% predictive accuracy for hospital assets — we achieve 72.1%.
Similar studies using IoT telemetry (e.g., OnWatch) depend on real-time data; our framework achieves competitive accuracy using static SR data, making it more scalable.

## Slide 27: Summary Result
This study presents a fully data-driven predictive maintenance
framework combining statistical reliability and machine
learning for healthcare imaging systems.
Using real operational data, it demonstrates how MTBF and TTR analytics can drive practical improvements in maintenance planning.
The system achieves 72.1% accuracy, identifies high-risk assets,
and generates 60-day PM schedules. Beyond technical contribution.
It provides a blueprint for digital transformation of service operations in emerging markets, enabling sustainable,patient-centric healthcare engineering.

## Slide 28: Conclusion & Future Work
Conclusion: MTBF pipeline and risk scoring provide actionable fleet prioritization; TTR models usable for ETA guidance.
Future: Add PM delay/QARA linkage, richer telemetry (usage hours/alerts), site logistics features, and spare SLA signals.
Modeling: Calibrate per‑modality models, quantile regression for SLA bands, and SHAP for explainability.
Delivery: Publish PowerBI/Tableau dashboard; schedule weekly refresh from ProServ360.
Notes: MTBF Summary Statistics: Systems with multiple failures: 135 ,Average MTBF (hours): 2482.49, Median MTBF (hours): 2101.19
Prediction Accuracy Assessment: Systems with assessable predictions: 121 Mean Absolute Error: 1885.07 hours Mean Prediction Accuracy: 72.14%, Future Failure Predictions: Systems with predictions: 145 High risk systems: 91 Medium risk systems: 16 Low risk systems: 38
Total systems analyzed: 157 Systems with predictable patterns: 121 Average prediction accuracy: 72.1%
RISK ASSESSMENT: High risk systems (immediate attention needed): 91 Medium risk systems (monitor closely): 16 Low risk systems (normal operation): 38 Results saved to: - MTBF_Analysis_Results.xlsx (Excel file with all analysis) - MTBF_Analysis_Visualizations.png (Charts and graphs)
MTBF ANALYSIS COMPLETE ================================================== Total systems analyzed: 157 Systems with predictable patterns: 121 Average prediction accuracy: 72.1% Average Suggested PM Interval: 82.7 days
RISK ASSESSMENT: High risk systems (immediate attention needed): 91 Medium risk systems (monitor closely): 16 Low risk systems (normal operation): 38 PM SCHEDULE SUMMARY: Systems due for PM within 7 days: 107 Systems due for PM within 30 days: 124 Recommended PM frequency range: 13.7 to 383.1 days Results saved to: - MTBF_Analysis_Results.xlsx (Excel file with all analysis) - MTBF_Analysis_Visualizations.png (Charts and graphs)
Available outlier removal methods: 1. iqr - Interquartile Range (recommended for skewed data) 2. zscore - Z-score (3 standard deviations) 3. quantile - 5th to 95th percentile 4. modality_specific - IQR per modality group ============================================================ DATA CLEANING AND OUTLIER REMOVAL ============================================================ Original data points: 1430 IQR Method: Q1=1.55, Q3=15.07, IQR=13.52 Bounds: [-18.73, 35.36] Removed 223 outliers (15.6% of data) Remaining data points: 1207 New TTR range: 0.02 to 35.16 hours New average TTR: 6.59 hours

## Slide 29: References
ProServ360 Service Logs (2023–2025), internal extracts.
Niyonambaza et al., Future Internet (2020) — IoT‑based PdM in hospitals.
scikit‑learn & xgboost docs; pandas/Matplotlib docs.
Kurose & Ross (for method presentation structure).